## Sheet1
Device | Test Method | Notes | Nominal Settings | Nominal Readings | Overloaded Settings | Overloaded Readings
CANBus | Power Off | Test Inputs are high impedance when board is powered down. Vary voltage 0V-10V on each pin. Verify <1mA current sink.
Power Supplies | Power On | Power the board by applying 36V to J12-11 & J12-12 and the Ground return to J12-5 & J12-6.
LED Red Driver | Non CAN Power Up | Power up machine in Non CAN mode. This driver should turn ON briefly and voltage should be read at J12-3. | Load=250Ohm | 3V +/- .5V
LED Green Driver | Non CAN Power Up | Power up machine in Non CAN mode. This driver should turn ON briefly and voltage should be read at J12-2. | Load=250Ohm | 3V +/- .5V
CAN Transciever | CAN Comms | If CAN heartbeats are working, this works OK.
Config Detect | CAN Data | Query CAN data for Inputs before and after toggling J12-7 | Open=Off/Ground=On
Enable Low Input | CAN Data | Query CAN data for Inputs before and after toggling J12-4 | Open=Off/Ground=On
Panel Switch Input | CAN Data | Query CAN data for Inputs before and after toggling J12-10 | Open=Off/B+=On
Pressure Switch Input | CAN Data | Query CAN data for Inputs before and after toggling J12-9 | Open=Off/Ground=On
Flush Switch Input | CAN Data | Query CAN data for Inputs before and after toggling J11-2 | Open=Off/Ground=On
Membrane Switch Input | Bed of Nails | Query CAN data for Inputs before and after toggling pin at J10-2. | Open=Off/Closed to J10-6=On
Membrane LED Low | Bed of Nails | Command LED states over CAN. Check pin state at J10-3. | Open=Off/Ground=On
Membrane LED Med | Bed of Nails | Command LED states over CAN. Check pin state at J10-4. | Open=Off/Ground=On
Membrane LED High | Bed of Nails | Command LED states over CAN. Check pin state at J10-5. | Open=Off/Ground=On
Membrane LED Status | Bed of Nails | Check status at J10-1 with and without Cell running nominally. | Open=Off/Ground=On
Pump Output | CAN Data | Load low-side connected to J12-8. Command Pump to Run. Verify PWM readings from CAN data and voltage to load. | V=10V, Load=36Ohm | Duty Cycle=27.8% | V=10V, Load=4.5Ohm | Pump Short Fault
Pump Current FB | CAN Data | Query CAN data for Pump current, verify within 5% of load current. |  | 1A |  | >85.4A
Peristaltic Pump Output | CAN Data | Load connected between J12-13 and J12-14. Command Pump to Run. Verify 3V between J12-13 and J12-14. | 3.3V | 12mA | 3.3V | 200mA
Peristaltic Pump Fault Detection | CAN Data | Open Circuit, check CAN error Code. Short Pump, check CAN error Code. | Open Load | Open Fault | see above | Shorted Fault
Cell Output | CAN Data | Command Cell to run. Current will be driven between J11-3/6 and J11-4/5. Verify Verify PWM readings from CAN data and voltage to load. | Target Current=5.0A, Load=3.6Ohm | PWM=50% | Target Current=5.0A, Load=0.5Ohm | PWM=5%
Cell Current FB | CAN Data | Query CAN data for Cell current, verify within 5% of load current. |  | 5.0A |  | >85.4A
SPI Flash | CAN Comms | Test with Parameter Write/Read of Build date and Board Serial Number.
RTC | CAN Comms | Command RTC Date/Time Set current GMT (ideally) or 01/01/2000 (backup). Power down board, pause, power up. Read RTC Date/Time from CAN and verify correct date and an appropriate amount of seconds have incremented.
Battery Voltage Reading | CAN Data | Query CAN data for water pump battery voltage, verify within 1% of actual voltage.
Note on B4: Meyer, Nathan: Power up with J12-7 tied to Ground to enter Non-CAN mode.
Note on B5: Meyer, Nathan: Power up with J12-7 tied to Ground to enter Non-CAN mode.
Note on B6: Meyer, Nathan: Power up with J12-7 floating to enable CAN mode for remaining tests.
Note on E19: Meyer, Nathan: DC current
Note on G18: Hawkins, Charles: 85.4A is trip current Overload can be set to 90A to get trip to happen.
Note on G19: Meyer, Nathan: DC Current
Note on G22: Hawkins, Charles: 85.4A is trip current Overload can be set to 90A to get trip to happen.